FeNOse NO Concentration Predictor

Project File & Architecture Report

Generated: 30 March 2026

# 1. Project Overview

The FeNOse project is a machine learning system that predicts exhaled nitric oxide (NO) concentration in parts per billion (ppb) from raw breathalyzer sensor data. NO concentration is a key clinical biomarker for airway inflammation and is routinely measured in asthma diagnosis and management.

The system uses a custom Multi-Layer Perceptron (MLP) trained on data from six FeNOse breathalyzer devices. Two model versions have been produced: v1 (feature selection + log-scaled output) and v2 (PCA dimensionality reduction + linear output). A Streamlit web application provides a point-and-click interface for uploading raw CSV files and viewing predictions in real time.

# 2. Project Location

All project files are stored in the user’s selected folder. The top-level structure is:

| File / Folder | Size | Description |
| --- | --- | --- |
| data_fenoze/model/ | ~470 KB | Trained model weights, preprocessing params, metrics, plots |
| data_fenoze/model_tester/ | ~21 KB | Streamlit web application for inference |
| data_fenoze/09/ | ~CSV | Raw measurements — Device 09 (original batch) |
| data_fenoze/09-2/ | ~CSV | Raw measurements — Device 09 (second batch, 475 files) |
| data_fenoze/18/ | ~CSV | Raw measurements — Device 18 (original batch) |
| data_fenoze/18-2/ | ~CSV | Raw measurements — Device 18 (second batch) |
| data_fenoze/27/ | ~CSV | Raw measurements — Device 27 (185 files) |
| data_fenoze/36/ | ~CSV | Raw measurements — Device 36 |
| data_fenoze/45/ | ~CSV | Raw measurements — Device 45 |
| data_fenoze/54/ | ~CSV | Raw measurements — Device 54 |
| data_fenoze/by_device/ | ~CSV | All CSVs reorganised by device serial (16 device sub-folders) |

# 3. Raw Data

## 3.1  Dataset Statistics

| Statistic | Value | Notes |
| --- | --- | --- |
| Total CSV files | 1,780 | Across all device folders |
| Unique devices | 6 (IDs: 9, 18, 27, 36, 45, 54) | Plus 16 sub-device sessions in by_device/ |
| Concentration levels | 7 | 0, 5, 10, 20, 50, 75, 100 ppb |
| 0 ppb files | 145 | Zero / clean-air reference |
| 5 ppb files | 145 |  |
| 10 ppb files | 136 |  |
| 20 ppb files | 148 |  |
| 50 ppb files | 142 |  |
| 75 ppb files | 142 |  |
| 100 ppb files | 130 | Maximum calibration concentration |

## 3.2  File Naming Convention

Each CSV filename encodes key metadata. Example:

ABGv5MUXOff-75ppb-3AH-75ccm-Maxwell-Q126_b4d90388_0000000018-0926-asu-nz_..._20260316-2007.csv

| ABGv5MUXOff | Firmware / hardware configuration identifier |
| --- | --- |

| 75ppb | Actual NO concentration used in this measurement |
| --- | --- |

| 3AH | Breathing protocol (3 seconds hold) |
| --- | --- |

| 75ccm | Flow rate in cubic centimetres per minute |
| --- | --- |

| 0000000018 | Device serial number (device ID 18) |
| --- | --- |

| 20260316-2007 | Date and time of measurement (YYYYMMDD-HHMM) |
| --- | --- |

## 3.3  CSV File Structure

Each CSV contains time-series rows with the following key columns:

- event_name — phase label: AmbientSamplingRFC, FeNOMeasurement, FeNOWindow

- A1 – H8 — 64 sensor channels (8 rows × 8 columns on the sensor array)

- AQT0, AQH0, AQP0 — ambient temperature, humidity, pressure

# 4. Feature Extraction

Feature extraction is implemented in model/fenose_predict.py and replicated in model_tester/app.py. It operates on the three measurement phases within each CSV:

## 4.1  Phase Segmentation

- AmbientSamplingRFC — baseline ambient air reading (used as reference)

- FeNOMeasurement — the exhaled breath sample (primary signal)

- FeNOWindow — optional window phase (used if present)

## 4.2  Feature Engineering (265 features total)

| Feature Group | Count | Description |
| --- | --- | --- |
| Delta (d_*) | 64 | FeNO mean minus ambient mean per sensor (feno_mean − amb_mean) |
| Normalised delta (nd_*) | 64 | Delta divided by absolute ambient mean: delta / (¦amb_mean¦ + 1e-6) |
| FeNO std (fs_*) | 64 | Standard deviation of FeNOMeasurement phase per sensor |
| Window delta (wd_*) | 64 | FeNOWindow mean minus ambient mean (zeros if phase absent) |
| Env. variables | 3 | Mean ambient temperature (AQT0), humidity (AQH0), pressure (AQP0) |
| Global stats | 6 | delta mean/max/min/std and normalised delta mean/std |

# 5. Model Architecture

## 5.1  Model v1 (LODO-validated)

v1 is the cross-validated model trained using Leave-One-Device-Out strategy.

| Input features | 265 extracted → 80 selected (top_idx feature selection) |
| --- | --- |

| Preprocessing | Standardisation: (X − mean) / std |
| --- | --- |

| Architecture | MLP: 80 → 64 → 32 → 1 (ReLU activations) |
| --- | --- |

| Output | Log-ppb space: prediction = expm1(clip(raw, 0, 10)) |
| --- | --- |

| Weights file | model/fenose_mlp_weights.npz (59 KB) |
| --- | --- |

| Preprocessing file | model/fenose_preprocessing.npz (14 KB) |
| --- | --- |

## 5.2  Model v2 (PCA + linear output)

v2 uses PCA for dimensionality reduction and a linear output scaled by y_max.

| Input features | 265 extracted → 265 (all valid via good_mask) |
| --- | --- |

| Preprocessing | Standardisation → PCA reduction to 50 components |
| --- | --- |

| Architecture | MLP: 50 → 48 → 24 → 1 (ReLU activations) |
| --- | --- |

| Output | Linear: prediction = clip(raw × 100, 0, 100) ppb |
| --- | --- |

| Weights file | model/fenose_mlp_weights_v2.npz (30 KB) |
| --- | --- |

| Preprocessing file | model/fenose_preprocessing_v2.npz (120 KB, includes V_pca) |
| --- | --- |

# 6. Model Performance

## 6.1  v1 — Leave-One-Device-Out Cross-Validation

| Device | MAE (ppb) | RMSE (ppb) |
| --- | --- | --- |
| Device 9 | 1.58 | 2.66 |
| Device 18 | 1.03 | 1.67 |
| Device 27 | 1.67 | 2.87 |
| Device 36 | 1.15 | 1.62 |
| Device 45 | 1.52 | 2.48 |
| Device 54 | 5.16 | 8.07 |
| Average | 2.02 | 3.23 |

Device 54 shows higher error, likely due to fewer training examples when that device is held out. The Ridge baseline (for comparison) achieved MAE = 36.96 ppb, confirming the MLP is highly effective for this task.

## 6.2  v2 — 80/20 Stratified Split

| MAE (ppb) | RMSE (ppb) | R² | Test Samples |
| --- | --- | --- | --- |
| 5.04 | 6.86 | 0.9633 | 56 |

## 6.3  v2 — Per-Concentration MAE

| 0 ppb | 5 ppb | 10 ppb | 20 ppb | 50 ppb | 75 ppb | 100 ppb |
| --- | --- | --- | --- | --- | --- | --- |
| 6.24 | 8.19 | 4.10 | 3.66 | 4.20 | 3.60 | 5.30 |

# 7. Training Details

## 7.1  Training Data

- Real samples: 278 (original measurements across 6 devices)

- Synthetic samples: 1,010 (generated via data augmentation)

- Total training set: 1,288 samples

- Features used in v1: 80 (selected from 265 total)

## 7.2  Data Augmentation Techniques

- Gaussian jitter — adds small random noise to sensor readings to improve robustness

- Concentration interpolation — generates synthetic samples between known concentration levels

- Cross-device mixing — blends features from different devices to improve generalisation

## 7.3  Validation Strategy

- v1: Leave-One-Device-Out (LODO) — train on 5 devices, test on the held-out device, repeat for all 6

- v2: 80/20 stratified random split — 222 real + 1,359 augmented training, 56 real test

# 8. Inference Pipeline

The inference pipeline takes a single raw CSV file and returns a predicted NO concentration in ppb. It is implemented in two places:

- model/fenose_predict.py — standalone command-line script

- model_tester/app.py — full web application with UI

## 8.1  Step-by-Step Pipeline

| Step | Stage | Detail |
| --- | --- | --- |
| 1 | Load CSV | Read raw CSV with pandas; filter rows by event_name phase |
| 2 | Phase split | Separate AmbientSamplingRFC, FeNOMeasurement, FeNOWindow rows |
| 3 | Compute delta | delta = feno_mean − amb_mean for each of 64 sensors |
| 4 | Normalise | norm_d = delta / (¦amb_mean¦ + 1e-6) per sensor |
| 5 | Build feature vec | Concatenate d_*, nd_*, fs_*, wd_*, env_*, global stats → 265-d vector |
| 6a | v1 — Select | Apply top_idx to select 80 features; standardise with saved mean/std |
| 6b | v2 — PCA | Standardise all 265 features; project through V_pca → 50 components |
| 7 | MLP forward pass | Layer 1 (ReLU) → Layer 2 (ReLU) → Layer 3 (linear) |
| 8a | v1 — Decode | ppb = expm1(clip(raw, 0, 10)) [log-space output] |
| 8b | v2 — Scale | ppb = clip(raw × 100, 0, 100) [linear output] |

# 9. Web Application (model_tester/)

## 9.1  File List

| File / Folder | Size | Description |
| --- | --- | --- |
| model_tester/app.py | 19 KB | Main Streamlit app — UI, file upload, prediction, charts |
| model_tester/requirements.txt | 0.1 KB | Python dependencies: streamlit, pandas, numpy, plotly |
| model_tester/run.sh | 0.5 KB | Mac / Linux launch script |
| model_tester/run.bat | 0.5 KB | Windows launch script |

## 9.2  How to Launch

Prerequisites: Python 3.8+ and pip installed.

Mac / Linux

cd path/to/data_fenoze/model_tester

pip3 install streamlit pandas numpy plotly

python3 -m streamlit run app.py --server.address 0.0.0.0 --server.port 8501

Windows

cd path\to\data_fenoze\model_tester

pip install streamlit pandas numpy plotly

python -m streamlit run app.py --server.address 0.0.0.0 --server.port 8501

The app will open at http://localhost:8501. Other devices on the same Wi-Fi network can access it via the host machine’s IP address.

## 9.3  App Features

- Model version selector (v1 / v2) in the sidebar

- Multi-file upload — drag and drop one or more raw CSV files

- Automatic actual concentration parsing from the filename (e.g. 75ppb)

- Single-file view: predicted ppb, actual ppb, absolute error, clinical range label

- Multi-file view: MAE, RMSE, R², scatter plot, error-by-concentration bar chart, results table

- Download predictions as CSV

# 10. Model Files (model/)

| File / Folder | Size | Description |
| --- | --- | --- |
| fenose_mlp_weights.npz | 59 KB | v1 MLP weight matrices W1, b1, W2, b2, W3, b3 |
| fenose_preprocessing.npz | 14 KB | v1 preprocessing: feat_cols, top_idx, scaler_mean, scaler_std |
| fenose_mlp_weights_v2.npz | 30 KB | v2 MLP weight matrices (50→48→24→1) |
| fenose_preprocessing_v2.npz | 120 KB | v2: feat_cols, good_mask, feat_mean, feat_std, V_pca (265→50), y_max |
| fenose_predict.py | 2.8 KB | Standalone command-line inference script |
| metrics.json | 0.9 KB | v1 LODO cross-validation results per device |
| validation_80_20_metrics.json | 0.3 KB | v2 80/20 split results: MAE, RMSE, R², per-concentration MAE |
| model_evaluation.png | 264 KB | v1 model evaluation plot |
| validation_80_20.png | 241 KB | v2 actual vs predicted scatter plot |

# 11. Clinical Reference

The FeNO (Fractional exhaled Nitric Oxide) test measures airway inflammation. Standard clinical thresholds:

| Range | Classification | Clinical Interpretation |
| --- | --- | --- |
| < 25 ppb | Normal | Low likelihood of eosinophilic airway inflammation |
| 25–50 ppb | Borderline | Possible airway inflammation; interpret with clinical context |
| ≥ 50 ppb | Elevated | Significant eosinophilic airway inflammation likely |

# 12. Quick Reference Summary

| Total raw CSV files | 1,780 across 8 device folders + by_device/ |
| --- | --- |

| Sensor channels | 64 (A1–H8 grid) |
| --- | --- |

| Features per sample | 265 (delta, normalised delta, std, window delta, env, global stats) |
| --- | --- |

| v1 model | 80 selected features → MLP 80→64→32→1, log-ppb output |
| --- | --- |

| v2 model | 265 features → PCA-50 → MLP 50→48→24→1, linear output |
| --- | --- |

| v1 cross-val MAE | 2.02 ppb (Leave-One-Device-Out) |
| --- | --- |

| v2 test MAE | 5.04 ppb (80/20 split, R² = 0.9633) |
| --- | --- |

| Inference script | model/fenose_predict.py |
| --- | --- |

| Web app | model_tester/app.py — run with: python3 -m streamlit run app.py |
| --- | --- |

FeNOse Project — Confidential  |  Generated 30 March 2026